## Toluene R classic
Examples of QC Charts that are build on the model and in assumption of the Gaussian (Normal) distribution of data
The charts below are so called "moving QC charts", since the current performance is analyzed against the most recent historical data. In this case the last 20 sets of LCS points are plotted against the historical warning and critical limits
[6 empty] | Accuracy |  | X-Chart |  |  |  |  | Precision | R-Chart
Run | Date | Batch | LCS1 | LCS2 | Average | X Chart | Historical limits |  |  |  | Range | LCS Range | R-Chart | Historical limits |  |  | Accuracy | Important constants, and results
|  |  | Recovery | Recovery | Recovery | LCL | LWL | Mean | UWL | UCL |  | Abs | Mean | UWL | UCL
20 | 2021-03-10 | 2022-1000 | 1 | 1.06 | 1.03 | 0.92389 | 0.96503 | 1.0781 | 1.1912 | 1.2323 |  | 0.06 | 0.024762 | 0.062177 | 0.080884
19 | 2020-12-16 | 2022-999 | 0.95 | 1.02 | 0.985 | 0.92389 | 0.96503 | 1.0781 | 1.1912 | 1.2323 |  | 0.07 | 0.024762 | 0.062177 | 0.080884
18 | 2020-12-11 | 2022-998 | 0.97 | 0.94 | 0.955 | 0.92389 | 0.96503 | 1.0781 | 1.1912 | 1.2323 |  | 0.03 | 0.024762 | 0.062177 | 0.080884 | K value @ 95% | 2.086 | Student's coefficient for 95% confidence. For Practical application, it is often rouded to 2.0
17 | 2020-10-07 | 2022-997 | 0.99 | 0.96 | 0.975 | 0.92389 | 0.96503 | 1.0781 | 1.1912 | 1.2323 |  | 0.03 | 0.024762 | 0.062177 | 0.080884 | K value @ 99% | 2.845 | Student's coefficient for 99% confidence. For practical application, it is often rounded up to 3.0
16 | 2020-09-14 | 2022-996 | 0.96 | 0.98 | 0.97 | 0.92389 | 0.96503 | 1.0781 | 1.1912 | 1.2323 |  | 0.02 | 0.024762 | 0.062177 | 0.080884 | MU | 0.054201 | Measurement Uncertainty calculated from historical data
15 | 2020-09-14 | 2022-995 | 0.89 | 0.87 | 0.88 | 0.92389 | 0.96503 | 1.0781 | 1.1912 | 1.2323 |  | 0.02 | 0.024762 | 0.062177 | 0.080884 | Extended MU @ 95% | 0.11306 | Measurement Uncertainty (Error) calculated at 95% confidence interval
14 | 2020-08-19 | 2022-994 | 1.02 | 0.98 | 1 | 0.92389 | 0.96503 | 1.0781 | 1.1912 | 1.2323 |  | 0.04 | 0.024762 | 0.062177 | 0.080884 | Acceptance Criteria MAX | 1.2323 | =UCL
13 | 2020-07-28 | 2022-993 | 1 | 0.99 | 0.995 | 0.92389 | 0.96503 | 1.0781 | 1.1912 | 1.2323 |  | 0.01 | 0.024762 | 0.062177 | 0.080884 | Acceptance Criteria MIN | 0.92389 | =LCL
12 | 2020-07-13 | 2022-992 | 0.98 | 0.96 | 0.97 | 0.92389 | 0.96503 | 1.0781 | 1.1912 | 1.2323 |  | 0.02 | 0.024762 | 0.062177 | 0.080884
11 | 2020-03-23 | 2022-991 | 0.93 | 0.95 | 0.94 | 0.92389 | 0.96503 | 1.0781 | 1.1912 | 1.2323 |  | 0.02 | 0.024762 | 0.062177 | 0.080884
10 | 2020-02-13 | 2022-990 | 1.14 | 1.15 | 1.145 | 0.92389 | 0.96503 | 1.0781 | 1.1912 | 1.2323 |  | 0.01 | 0.024762 | 0.062177 | 0.080884
9 | 2019-12-27 | 2022-989 | 0.95 | 1.01 | 0.98 | 0.92389 | 0.96503 | 1.0781 | 1.1912 | 1.2323 |  | 0.06 | 0.024762 | 0.062177 | 0.080884
8 | 2019-11-21 | 2022-988 | 1.18 | 1.13 | 1.155 | 0.92389 | 0.96503 | 1.0781 | 1.1912 | 1.2323 |  | 0.05 | 0.024762 | 0.062177 | 0.080884
7 | 2019-06-11 | 2022-987 | 1.17 | 1.03 | 1.1 | 0.92389 | 0.96503 | 1.0781 | 1.1912 | 1.2323 |  | 0.14 | 0.024762 | 0.062177 | 0.080884 |  | Precision | Important constants, and results
6 | 2019-06-10 | 2022-986 | 1 | 1.02 | 1.01 | 0.92389 | 0.96503 | 1.0781 | 1.1912 | 1.2323 |  | 0.02 | 0.024762 | 0.062177 | 0.080884
5 | 2019-05-24 | 2022-985 | 1.04 | 1.02 | 1.03 | 0.92389 | 0.96503 | 1.0781 | 1.1912 | 1.2323 |  | 0.02 | 0.024762 | 0.062177 | 0.080884
4 | 2019-05-13 | 2022-984 | 0.99 | 0.95 | 0.97 | 0.92389 | 0.96503 | 1.0781 | 1.1912 | 1.2323 |  | 0.04 | 0.024762 | 0.062177 | 0.080884
3 | 2019-03-19 | 2022-983 | 1.03 | 0.99 | 1.01 | 0.92389 | 0.96503 | 1.0781 | 1.1912 | 1.2323 |  | 0.04 | 0.024762 | 0.062177 | 0.080884 | K value @ 95% | - | Student's coefficient for 95% confidence
2 | 2019-01-22 | 2022-982 | 1.06 | 1.04 | 1.05 | 0.92389 | 0.96503 | 1.0781 | 1.1912 | 1.2323 |  | 0.02 | 0.024762 | 0.062177 | 0.080884 | K value @ 99% | - | Student's coefficient for 99% confidence
1 | 2019-01-14 | 2022-981 | 1.05 | 1.03 | 1.04 | 0.92389 | 0.96503 | 1.0781 | 1.1912 | 1.2323 |  | 0.02 | 0.024762 | 0.062177 | 0.080884 | MU | - | Measurement Uncertainty calculated from historical data
0 | 2018-11-26 | 2022-980 | 1.1 | 1.09 | 1.095 | 0.92389 | 0.96503 | 1.0781 | 1.1912 | 1.2323 |  | 0.01 | 0.024762 | 0.062177 | 0.080884 | Extended MU @ 95% | - | Measurement Uncertainty (Error) calculated at 95% confidence interval
-1 | 2018-09-18 | 2022-979 | 1.01 | 1.04 | 1.025 | [6 empty] | 0.03 |  |  |  | Acceptance Criteria MAX | 0.080884 | =UCL
-2 | 2018-08-23 | 2022-978 | 1.09 | 1.07 | 1.08 | [6 empty] | 0.02 |  |  |  | Acceptance Criteria MIN | - | -not regulated for chemical; analysis
-3 | 2018-08-06 | 2022-977 | 1.05 | 1.05 | 1.05 | [6 empty] | 0 |  |  |  | d2 | 1.1284 | - Constants that are used to calculate D4
-4 | 2018-07-14 | 2022-976 | 1.02 | 1 | 1.01 | [6 empty] | 0.02 |  |  |  | d3 | 0.8525 | - Constants that are used to calculate D4
-5 | 2018-02-28 | 2022-975 | 1.14 | 1.06 | 1.1 | [6 empty] | 0.08 |  |  |  | D4 @ 95% | 2.511 | =1+2*(d3/d2) | Constant that is used to calculate UWL as UWL=D4(@95%)*Rmedian
-6 | 2018-01-12 | 2022-974 | 1.07 | 1.08 | 1.075 | [6 empty] | 0.01 |  |  |  | D4 @ 99% | 3.2665 | =1+3*(d3/d2) | Constant that is used to calculate UCL as UCL=D4(@99%)*Rmedian
-7 | 2018-01-06 | 2022-973 | 1.02 | 1.01 | 1.015 | [6 empty] | 0.01
-8 | 2017-12-26 | 2022-972 | 1.02 | 0.97 | 0.995 | [6 empty] | 0.05
-9 | 2017-12-19 | 2022-971 | 1.07 | 1.05 | 1.06 | [6 empty] | 0.02
-10 | 2017-09-25 | 2022-970 | 1.17 | 1.17 | 1.17 | [6 empty] | 0
-11 | 2017-09-22 | 2022-969 | 1.12 | 1.11 | 1.115 | [6 empty] | 0.01
-12 | 2017-07-25 | 2022-968 | 1.13 | 1.18 | 1.155 | [6 empty] | 0.05
-13 | 2017-07-25 | 2022-967 | 1.1 | 1.1 | 1.1 | [6 empty] | 0
-14 | 2017-07-25 | 2022-966 | 1.07 | 1.06 | 1.065 | [6 empty] | 0.01
-15 | 2017-07-19 | 2022-965 | 1.04 | 1.04 | 1.04 | [6 empty] | 0
-16 | 2017-07-11 | 2022-964 | 1.05 | 1.05 | 1.05 | [6 empty] | 0
-17 | 2017-07-08 | 2022-963 | 1.07 | 1.13 | 1.1 | [6 empty] | 0.06
-18 | 2017-07-07 | 2022-962 | 1.22 | 1.11 | 1.165 | [6 empty] | 0.11
-19 | 2017-06-16 | 2022-961 | 1.05 | 1.02 | 1.035 | [6 empty] | 0.03
-20 | 2017-06-05 | 2022-960 | 1.14 | 1.14 | 1.14 | [6 empty] | 0
-21 | 2017-03-26 | 2022-959 | 0.96 | 0.99 | 0.975 | [6 empty] | 0.03
-22 | 2017-03-25 | 2022-958 | 0.96 | 0.99 | 0.975 | [6 empty] | 0.03
-23 | 2017-01-24 | 2022-957 | 1.05 | 1 | 1.025 | [6 empty] | 0.05
-24 | 2017-01-17 | 2022-956 | 1.1 | 1.11 | 1.105 | [6 empty] | 0.01
-25 | 2016-10-11 | 2022-955 | 1 | 1.04 | 1.02 | [6 empty] | 0.04
-26 | 2016-09-22 | 2022-954 | 0.99 | 0.98 | 0.985 | [6 empty] | 0.01
-27 | 2016-08-31 | 2022-953 | 0.99 | 1.03 | 1.01 | [6 empty] | 0.04
-28 | 2016-07-18 | 2022-952 | 0.99 | 0.97 | 0.98 | [6 empty] | 0.02
-29 | 2016-05-02 | 2022-951 | 0.97 | 0.89 | 0.93 | [6 empty] | 0.08
-30 | 2016-05-02 | 2022-950 | 0.88 | 0.86 | 0.87 | [6 empty] | 0.02
-31 | 2016-04-21 | 2022-949 | 1.1 | 1.01 | 1.055 | [6 empty] | 0.09
-32 | 2016-02-10 | 2022-948 | 1.02 | 1.02 | 1.02 | [6 empty] | 0
-33 | 2016-02-10 | 2022-947 | 1.02 | 1.02 | 1.02 | [6 empty] | 0
-34 | 2016-02-10 | 2022-946 | 0.99 | 1.01 | 1 | [6 empty] | 0.02
-35 | 2016-02-10 | 2022-945 | 1.03 | 1.01 | 1.02 | [6 empty] | 0.02
-36 | 2016-02-10 | 2022-944 | 0.97 | 1.01 | 0.99 | [6 empty] | 0.04
-37 | 2016-02-10 | 2022-943 | 0.96 | 1.01 | 0.985 | [6 empty] | 0.05
-38 | 2016-02-10 | 2022-942 | 1 | 1 | 1 | [6 empty] | 0
-39 | 2016-02-10 | 2022-941 | 1.03 | 0.99 | 1.01 | [6 empty] | 0.04
-40 | 2016-01-24 | 2022-940 | 0.91 | 0.99 | 0.95 | [6 empty] | 0.08
-41 | 2016-01-21 | 2022-939 | 1.02 | 1.06 | 1.04 | [6 empty] | 0.04
-42 | 2016-01-05 | 2022-938 | 1.03 | 0.98 | 1.005 | [6 empty] | 0.05
-43 | 2015-12-16 | 2022-937 | 0.87 | 0.94 | 0.905 | [6 empty] | 0.07
-44 | 2015-11-25 | 2022-936 | 0.92 | 0.96 | 0.94 | [6 empty] | 0.04
-45 | 2015-11-09 | 2022-935 | 1.03 | 1.08 | 1.055 | [6 empty] | 0.05
-46 | 2015-10-20 | 2022-934 | 1.03 | 0.95 | 0.99 | [6 empty] | 0.08
-47 | 2015-10-12 | 2022-933 | 1.01 | 0.95 | 0.98 | [6 empty] | 0.06
-48 | 2015-08-11 | 2022-932 | 1.09 | 0.86 | 0.975 | [6 empty] | 0.23
-49 | 2015-07-24 | 2022-931 | 0.98 | 0.96 | 0.97 | [6 empty] | 0.02
-50 | 2015-05-29 | 2022-930 | 0.98 | 1.01 | 0.995 | [6 empty] | 0.03
-51 | 2015-05-04 | 2022-929 | 1.02 | 1.09 | 1.055 | [6 empty] | 0.07
-52 | 2015-04-17 | 2022-928 | 1.04 | 1.04 | 1.04 | [6 empty] | 0
-53 | 2015-04-05 | 2022-927 | 1 | 1 | 1 | [6 empty] | 0
-54 | 2015-04-04 | 2022-926 | 1.04 | 1.04 | 1.04 | [6 empty] | 0
-55 | 2015-03-02 | 2022-925 | 1 | 1.03 | 1.015 | [6 empty] | 0.03
-56 | 2015-03-02 | 2022-924 | 0.95 | 0.97 | 0.96 | [6 empty] | 0.02
-57 | 2015-03-02 | 2022-923 | 0.99 | 0.97 | 0.98 | [6 empty] | 0.02
-58 | 2015-01-09 | 2022-922 | 1.03 | 1.02 | 1.025 | [6 empty] | 0.01
-59 | 2014-10-25 | 2022-921 | 0.99 | 1 | 0.995 | [6 empty] | 0.01
-60 | 2014-07-24 | 2022-920 | 1.01 | 0.99 | 1 | [6 empty] | 0.02
-61 | 2014-05-27 | 2022-919 | 1.03 | 1.03 | 1.03 | [6 empty] | 0
-62 | 2014-05-09 | 2022-918 | 0.8 | 0.84 | 0.82 | [6 empty] | 0.04
-63 | 2014-04-28 | 2022-917 | 0.95 | 0.95 | 0.95 | [6 empty] | 0
-64 | 2014-03-27 | 2022-916 | 1 | 0.96 | 0.98 | [6 empty] | 0.04
-65 | 2014-03-04 | 2022-915 | 1.03 | 1.06 | 1.045 | [6 empty] | 0.03
-66 | 2014-02-27 | 2022-914 | 1.07 | 1.05 | 1.06 | [6 empty] | 0.02
-67 | 2013-12-19 | 2022-913 | 1.15 | 1.1 | 1.125 | [6 empty] | 0.05
-68 | 2013-12-09 | 2022-912 | 0.97 | 1.02 | 0.995 | [6 empty] | 0.05
-69 | 2013-11-27 | 2022-911 | 1.04 | 1.01 | 1.025 | [6 empty] | 0.03
-70 | 2013-08-23 | 2022-910 | 0.93 | 1 | 0.965 | [6 empty] | 0.07
-71 | 2013-07-31 | 2022-909 | 0.88 | 0.87 | 0.875 | [6 empty] | 0.01
-72 | 2013-06-17 | 2022-908 | 0.89 | 0.85 | 0.87 | [6 empty] | 0.04
-73 | 2013-02-28 | 2022-907 | 0.87 | 0.88 | 0.875 | [6 empty] | 0.01
-74 | 2012-05-04 | 2022-906 | 0.92 | 0.96 | 0.94 | [6 empty] | 0.04
-75 | 2011-11-24 | 2022-905 | 1.05 | 1.09 | 1.07 | [6 empty] | 0.04
-76 | 2011-08-16 | 2022-904 | 1.08 | 1.08 | 1.08 | [6 empty] | 0
-77 | 2011-07-29 | 2022-903 | 1.04 | 1.02 | 1.03 | [6 empty] | 0.02
-78 | 2011-07-04 | 2022-902 | 1.08 | 1.08 | 1.08 | [6 empty] | 0
-79 | 2011-06-07 | 2022-901 | 1.06 | 1.03 | 1.045 | [6 empty] | 0.03
-80 | 2011-05-10 | 2022-900 | 1.06 | 1.06 | 1.06 | [6 empty] | 0
-81 | 2011-01-13 | 2022-899 | 1.0964 | 1.1205 | 1.1084 | [6 empty] | 0.024096
-82 | 2011-01-13 | 2022-898 | 1.0843 | 1.0843 | 1.0843 | [6 empty] | 0
-83 | 2011-01-13 | 2022-897 | 1.0361 | 1.0843 | 1.0602 | [6 empty] | 0.048193
-84 | 2011-01-13 | 2022-896 | 1 | 1.0602 | 1.0301 | [6 empty] | 0.060241
-85 | 2011-01-13 | 2022-895 | 1.0602 | 1.0602 | 1.0602 | [6 empty] | 0
-86 | 2011-01-13 | 2022-894 | 1.0361 | 1.0482 | 1.0422 | [6 empty] | 0.012048
-87 | 2011-01-13 | 2022-893 | 1.0361 | 1.0482 | 1.0422 | [6 empty] | 0.012048
-88 | 2011-01-13 | 2022-892 | 1 | 1.0482 | 1.0241 | [6 empty] | 0.048193
-89 | 2011-01-13 | 2022-891 | 1.0241 | 1.0361 | 1.0301 | [6 empty] | 0.012048
-90 | 2011-01-13 | 2022-890 | 1 | 1.012 | 1.006 | [6 empty] | 0.012048
-91 | 2011-01-03 | 2022-889 | 1.0644 | 1.125 | 1.0947 | [6 empty] | 0.060606
-92 | 2011-01-03 | 2022-888 | 1.0644 | 1.1212 | 1.0928 | [6 empty] | 0.056818
-93 | 2011-01-03 | 2022-887 | 1.1061 | 1.1174 | 1.1117 | [6 empty] | 0.011364
-94 | 2011-01-03 | 2022-886 | 1.1061 | 1.0606 | 1.0833 | [6 empty] | 0.045455
-95 | 2011-01-03 | 2022-885 | 1.072 | 1.0455 | 1.0587 | [6 empty] | 0.026515
-96 | 2011-01-03 | 2022-884 | 1.0644 | 1.0379 | 1.0511 | [6 empty] | 0.026515
-97 | 2011-01-03 | 2022-883 | 1.072 | 1.0303 | 1.0511 | [6 empty] | 0.041667
-98 | 2011-01-03 | 2022-882 | 0.98106 | 1.0227 | 1.0019 | [6 empty] | 0.041667
-99 | 2011-01-03 | 2022-881 | 1.0189 | 1 | 1.0095 | [6 empty] | 0.018939
-100 | 2011-01-03 | 2022-880 | 0.98106 | 0.99621 | 0.98864 | [6 empty] | 0.015152
-101 | 2010-07-14 | 2022-879 | 0.98 | 0.95 | 0.965 | [6 empty] | 0.03
-102 | 2010-03-01 | 2022-878 | 0.98717 | 1.0523 | 1.0197 | [6 empty] | 0.065153
-103 | 2010-03-01 | 2022-877 | 0.94472 | 1.0375 | 0.99112 | [6 empty] | 0.092794
-104 | 2010-03-01 | 2022-876 | 0.93189 | 0.98815 | 0.96002 | [6 empty] | 0.056269
-105 | 2010-03-01 | 2022-875 | 0.98322 | 0.98026 | 0.98174 | [6 empty] | 0.0029615
-106 | 2010-03-01 | 2022-874 | 0.96446 | 0.95459 | 0.95953 | [6 empty] | 0.0098717
-107 | 2010-03-01 | 2022-873 | 0.94077 | 0.95459 | 0.94768 | [6 empty] | 0.01382
-108 | 2010-03-01 | 2022-872 | 0.93583 | 0.94472 | 0.94028 | [6 empty] | 0.0088845
-109 | 2010-03-01 | 2022-871 | 0.93386 | 0.94373 | 0.9388 | [6 empty] | 0.0098717
-110 | 2010-03-01 | 2022-870 | 1.0405 | 0.93978 | 0.99013 | [6 empty] | 0.10069
-111 | 2010-03-01 | 2022-869 | 0.93485 | 0.91017 | 0.92251 | [6 empty] | 0.024679
-112 | 2010-02-05 | 2022-868 | 1.12 | 1.02 | 1.07 | [6 empty] | 0.1
-113 | 2010-02-05 | 2022-867 | 0.98 | 0.99 | 0.985 | [6 empty] | 0.01
Formulas:
Date = [Date, row above]
Note on A29: Stepan: Run number where #20 corresponds to the newest set of data, and #1 is the oldest set that is included in the plot
Note on B29: Stepan: Date of the experiment
Note on C29: Stepan: Batch ID or Any reference ID. For example project name
Note on E29: Stepan: Relative recovery of the Laboratory Control Sample #1 (LCS1)
Note on F29: Stepan: Relative recovery of the Laboratory Control Sample #2 (LCS2)
Note on G29: Stepan: Average of 2 recoveries =(LCS1+LCS2)/2
Note on H51: Stepan: The LCL is calculated from historical data as LCL=Median-K(@99)*MU
Note on I51: Stepan: The LWL is calculated from historical data as LWL=Median-K(@95)*MU
Note on J51: Stepan: Median value of historical data
Note on K51: Stepan: The UWL is calculated from historical data as UWL=Median+K(@95)*MU
Note on L51: Stepan: The UCL is calculated from historical data as UCL=Median+K(@99)*MU
Note on Q51: Stepan: Median value that is based on historical data
Note on R51: Stepan: UWL is based on historical data and calculated as UWL=D4(@95)* Rmedian
Note on S51: Stepan: UCL is based on historical data and calculated as UCL=D4(@99)* Rmedian